# Transfusion Service Guidelines

# Division

# Location

# Address

| ##### Overview | In order to provide the safest blood components for patients, Vitalant has established the following guidelines for providing pretransfusion testing and related services. The Transfusion Service is an FDA registered and AABB and CLIA certified laboratory that provides suitable blood components for transfusion. The laboratory performs compatibility testing and related serological testing procedures in accordance with the current version of the AABB Standards for Blood Banks and Transfusion Services. |
| --- | --- |

| ##### Table of Contents | The following table lists the contents of these Guidelines: |
| --- | --- |

| Topic | See Page |
| --- | --- |
| Contact Information | 4 |
| Definitions | 4 |
| General Requirements | 5 |
| Transfusion Facility Responsibilities | 6 |
| Orders for Transfusion or Testing | 7 |
| Specimen Collection and Labeling | 7 |
| Second Sample for ABO/Rh Determination | 8 |
| Specimen Transport | 8 |
| Evaluation of Orders | 8 |
| Issue and Transport of Blood Components | 9 |
| Blood Component Storage | 10 |
| Compatible Blood Components for Transfusion | 11 |
| Incompatible Red Cell Components for Transfusion | 11 |
| Blood Administration | 12-13 |
| Reflex Testing | 14 |
| Critical Results | 15 |
| Type and Screen | 16 |
| Emergency Requirement for Blood | 16 |
| Adverse Reactions | 17-22 |
| Signs and Symptoms of Transfusion Reactions | 17-20 |
| Process for Immediate Adverse Reaction | 21-22 |
| Process for Other Adverse Reactions | 22 |

| ##### Contact Information | Transfusion Service Contacts - Supervisor: - Manager: - Laboratory Phone: - Fax Number: - Director of Technical Services: - Medical Director: - Hospital Services: - Courier: - CLIA Number: - Other (specify) : - Other (specify) : |
| --- | --- |

| ##### Definitions | Transfusion Service (TS): The laboratory who performs pretransfusion testing and prepares and provides compatible blood components for patient transfusion. Transfusion Facility (TF): The hospital, clinic, or other health care facility who is responsible for the ordering and administration of blood components to the patient. Blood Supplier: The Blood Center that collects and delivers blood components to the Transfusion Service. Must: This word is used in this document to indicate a mandatory statement. Should: This word is used in this document to indicate a recommendation. Critical results (also known as alert or panic values), are laboratory results that indicate a need to notify the clinical staff of a change in the patient’s status or a possible life-threatening situation. |
| --- | --- |

| ##### General Requirements | These Requirements are based on the AABB Standards for Blood Banks and Transfusion Services and meet the requirements of other relevant regulatory agencies. The Transfusion Facility must have policies and procedures related to blood transfusion that are in compliance with these Standards - All manual records and specimen labels must be completed in indelible ink. Alternatively, records may be maintained electronically. - Positive patient identification during specimen collection and blood administration is the most critical control point to ensure a safe transfusion. - A Blood Bank or Hospital armband must be placed on any patient receiving a transfusion. - Patient specimens must be labeled at the bedside immediately after collection. - All identifying information (patient’s name, Facility Patient ID number, and the Barcoded Blood Bank or Hospital armband number) must be identical on the armbands, specimen label, and order documents (Transfusion Services Order, TS 003 or other equivalent document). Orders or Specimens that do not meet this requirement will be rejected. The Transfusion Facility will be notified as soon as possible if this situation occurs. In this event, new orders may be required and/or the specimen may need to be re-collected. Specimen re-labeling is not allowed. - Any record relating to compatibility testing and transfusion processes, including administration must be maintained a minimum of 10 years. |
| --- | --- |

| ##### Transfusion Facility Responsibilities | The Transfusion Facility is responsible for developing and maintaining policies, processes, and validated procedures that provide instructions for transfusion activities performed at the site. These include: - Obtaining informed consent and the associated form - Collection of specimens - Administration of blood products: - Confirming identity of the documents attached to each blood unit and patient identity prior to transfusion - Medical supervision of the transfusion administration process - Maintenance of equipment owned by and used in in the Transfusion Facility. - Process for providing emergency release of blood products that aligns with emergency services provided by the Transfusion Service - Appropriate storage and handling of blood components with provisions for isolation/quarantine of unsuitable components if units are removed from the Transfusion Service Transport Container - Evaluating and approving deviations from SOP within the Transfusion Facility - Recognition and reporting of adverse reactions - Reporting adverse events, biological product deviations, and transfusion-related fatalities to the FDA/CBER - Reporting post transfusion infectious diseases, including physician/patient notification and a ‘Lookback’ procedure - Management of Recall or Withdrawal notices - A quality system description, policy or procedure - Regular review of blood product utilization. - Notification to Transfusion Service of major changes to procedures - Staff training and competency assessment |
| --- | --- |

| ##### Orders for Transfusion or Testing | - Complete the Transfusion Service supplied Transfusion Service Order, TS 003 or approved Transfusion Facility equivalent form for acceptance of patient specimen. The Transfusion Facility must retain the physician orders as part of their records. - The information below is required on the order. - Patient name (first and last) - At least 1 second unique identifier, such as hospital identification number - Patient Date of Birth - Ordering physician - Ordering Facility - Testing ordered - Blood component(s) ordered, including any special needs (e.g., irradiated) - Specimen Collection Date/Time (may be recorded on the specimen) - Phlebotomist ID (may be recorded on the specimen) If possible, Transfusion Facility should include: - History of transfusion and/or pregnancy - Diagnosis - Medication history - History of transfusion reactions or known antibodies - Pretransfusion criteria when a component is ordered - Date and time components are needed |
| --- | --- |

| ##### Specimen Collection and Labeling | - Collect specimens after appropriate identification of the patient from armbands attached to the patient. - Blood Bank armbands may be supplied by the Transfusion Service for use by the Facility, if requested. - Armband and specimen label must be identical and include: - Patient’s name (first and last) - Hospital identification number - The date of specimen collection must be provided on the order or on the specimen label. - The Transfusion Facility must be able to the identity the phlebotomist collecting each sample. - A minimum of one 5-7 mL EDTA specimen should be collected for each order. Additional specimen may be required for patients with antibodies – consult with the Transfusion Service to determine what is needed. - A second sample for ABO/Rh determination is required for compatibility testing when no previous blood type is on file. Refer to next block. - Specimens generally expire three (3) days after collection – a specimen collected on day one expires at midnight on day four. - Expiration of specimens from patients who have not been pregnant or transfused in the past three months may be extended to 45 days - Store specimens at 2-8C. |
| --- | --- |

| ##### Second Sample for ABO/Rh Determination | Two determinations of the recipient’s ABO/Rh are required prior to transfusion of type specific red cell components. The first determination shall be performed on a current sample and the second determination by one of the following methods. - Testing a second current sample. - Ideally, the samples are drawn by two people drawing two separate samples. - If staffing is prohibitive, one staff member may draw two separate samples, at separate times and must include a separate verification of patient identification. - A previously collected, EDTA sample, which is properly labeled, and no more than 7 days old, can be sent for testing in lieu of an additional draw. - The second sample must be labeled with patient’s full name and at least one additional unique identifier. - Comparison with previous records, which may include: - Previous history in the Transfusion Service LIS or manual records. - NOTE: If a second sample cannot be obtained, contact the TS. |
| --- | --- |

| ##### Specimen Transport | Specimens must arrive at the Transfusion Service as soon as possible, but no later than 24 hours after collection of the specimen. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1 | - Package and ship specimens appropriately. - Include the TRANSFUSION SERVICES ORDER, TS 003 inside the container, separate from the specimen(s). |
| 2 | If the sample is being shipped to the Transfusion Service, place a completed address label on the outside of the shipping box. |
| 3 | Contact a courier/transport service or Transfusion Services Representative to arrange for transport. |

| ##### Evaluation of Orders | - Each order received in the Transfusion Service is reviewed against defined criteria. - Orders or specimens that lack critical information will be rejected and need to be resubmitted. |
| --- | --- |

| ##### Issue and Transport of Blood Components | Blood units are packed in a sealed, temperature-validated transport container according to the type of component being issued. - Blood Component transport is arranged by the Transfusion Service prior to transfusion time requested. - Components requiring different storage temperatures are packed in different transport containers. Facility may store red blood cell components in the Transfusion Service sealed transport container for up to 21 hours. Units must be transfused or returned to the TS within 21 hours from the time of issue, indicated on the Compatibility Tag. - All other components must be transfused as soon as possible after receipt or stored appropriately. |
| --- | --- |

| If the Transfusing Facility | Then |
| --- | --- |
| Does not have approved storage equipment | - Transfusion of components must be initiated as soon as possible from opening the transport container. |
|  | - Components must be transfused prior to sample expiration indicated on the compatibility tag and the expiration date of the component. |
| Maintains approved blood storage equipment (See the section on Blood Component Storage) | - Transfer blood components to storage unit immediately after opening transport container. |
|  | - Components must be transfused prior to sample expiration indicated on the compatibility tag and the expiration date of the component. |
| Does not have return privileges | - Destroy unused blood components at the end of storage container expiration time following appropriate regulation for disposal of biohazardous material using universal precautions. |
| Has return privileges | - Transport container may be returned to Transfusion Service if container remains sealed and in accordance with local established process. |

| ##### Blood Component Storage | Store all blood components according to FDA and AABB requirements. |
| --- | --- |

| Component | Storage Temperature | Other Considerations |
| --- | --- | --- |
| Red Blood Cells or Whole Blood | 1-6C |  |
| Plasma | < -20C when frozen |  |
|  | 1-6C after thawing |  |
| Platelets | 20-24C | Maintain continuous gentle agitation during prolonged storage (more than 6 – 8 hours) |
| Cryoprecipitated AHF or Pooled Cryoprecipitated AHF (Cryo) | < -18C when frozen |  |
|  | 20-24C after thawing |  |
| Granulocyte Concentrate | 20-24C | Requires crossmatch, and 48 hour notification prior to need. |

NOTES: - The Transfusion Service must agree to allow the Transfusion Facility to store blood components before the Transfusion Facility begins storing components. - Storage units and blood component handling by the Transfusion Facility is subject to regular inspection. - Blood storage equipment must be dedicated to blood components ONLY. - NO food and/or chemicals or drugs may be stored in the equipment. - Components must be transfused within the expiration date of the product - If frozen components are thawed by the Transfusion Facility, the proper expiration time must be assigned when the unit is thawed.

| ##### Compatible Blood Components for Transfusion | The following tables, defined by type of component, show the appropriate donor unit ABO Group and Rh type based on the patient’s blood type. |
| --- | --- |

| RED BLOOD CELLS |  |
| --- | --- |
| Recipient’s ABO Group | Acceptable RBC ABO Group |
| O | O |
| A | A or O |
| B | B or O |
| AB | AB, A, B, or O |
| Recipient Rh Type | Acceptable RBC Type |
| D Positive | D Positive or Negative |
| D Negative | D Negative (refer to NOTE) |

| PLASMA: No Rh Requirement |  |
| --- | --- |
| Recipient’s ABO Group | Acceptable Component ABO Group |
| O | O, A, B, or AB |
| A | A or AB |
| B | B or AB |
| AB | AB |

PLATELETS: ABO/Rh is not generally a consideration for platelet components. Special consideration should be given to the following patient groups: - Frequently transfused patients - Patients under the age of 18 - Women of childbearing age (under age 50) – for Rh only

CRYOPRECIPITATE: No ABO/Rh specific criteria required for this component.

NOTE: For females under the age of 50 and males under the age of 18, Rh type compatible red cell and platelet components should be provided whenever possible. - RhIG therapy should be considered for Rh Negative women under 50 years old who have received Rh Positive platelet products. - Local policies may establish more restrictive criteria. NOTE: Low Titer Platelets and Group A Plasma may be used inter-changeably with Group AB platelets and plasma in certain locations.

| ##### Incompatible Red Cell Components for Transfusion | For incompatible crossmatched red cell products issued, it is laboratory policy (led by medical staff) that the products provided are the best possible for the patient, at the moment of issue. Depending upon the individual patient situation, this policy may mean the product is least incompatible or the best phenotype matched product. |
| --- | --- |

| ##### Blood Administration | - Blood administration is a critical aspect of patient care and must be performed by qualified staff. - A physician order and patient signed Consent Form must be present in the patient’s chart prior to beginning the transfusion. - Each component provided by the Transfusion Service has a Compatibility/Transfusion Record attached to the component. - At time of issue, the component label is verified to ensure it matches the attached Compatibility/Transfusion Record. - Immediately prior to transfusion, the Compatibility/Transfusion Record must be compared to the component label and the patient’s Barcoded Hospital or Blood Bank armband, to ensure all information matches. - This must be performed by two individuals or one person using an automated verification system. - If the information does not match, the facility must contact the Transfusion Service and return the unit for appropriate resolution of the problem. - The Compatibility/Transfusion Record must remain attached to the donor unit until completion of the transfusion. - During the transfusion, the patient must be monitored for signs and symptoms of transfusion reactions, including monitoring of vitals. This is especially important during the first 15-30 minutes of the transfusion. - If a transfusion reaction is detected during or subsequent to the transfusion, stop the transfusion, but do not disconnect the tubing. Contact the Transfusion Service and the transfusing physician immediately - Upon completion of transfusion, Transfusion Facility personnel must take the patient’s vitals, and complete the Compatibility/Transfusion Record and place in patient’s chart. - Properly discard the empty blood product bag and tubing in a biohazard container. - A copy of the completed Compatibility/Transfusion Record may be sent to the Transfusion Service to provide documentation of unit disposition. - If transfusion was provided to an outpatient, provide specific written instructions concerning possible adverse events, including emergency medical care contacts. - Transfusion documentation retained at the Transfusion Facility must include the following (at a minimum). - Transfusion order - Recipient consent - Component name - Donation Identification Number (DIN) - Date and time of transfusion - Pre- and posttransfusion vital signs - Volume transfused - Identification of the transfusionist - Transfusion-related adverse events, if applicable |
| --- | --- |

Continued on next page

Blood Administration (continued)

|  | The following table presents general transfusion information by blood component: |
| --- | --- |

| Criteria | Red Blood Cells | Platelets/Plasma | Granulocytes | Cryoprecipitate |
| --- | --- | --- | --- | --- |
| Venous Access | 16-22 gauge | 22 gauge or larger | 22 gauge or larger | 22 gauge or larger |
| Routine in-line filter | Particulate filter 170 – 260 microns |  |  |  |
| Leukocyte reduction filter | May be used NOTE: Do not use on pre-storage leukocytes reduced components. |  | Do not use |  |
| Microaggregate filter | May be used |  |  |  |
| Product volume | ~ 300 mL | 150-270 mL | 200-300 mL | ~ 100 mL |
| Maximum time for transfusion | 4 hours |  |  | As rapidly as tolerated |
| Avg. time for routine transfusion | 1-4 hours | 30 min-1 hour | 2-4 hours |  |
| Slow rate | 2 mL per minute |  | 1 mL per minute |  |
| Avg. rate | 2-5 mL/minute | 4-10 mL/minute or as tolerated | Do not give rapidly |  |

|  | NOTES: - Normal Saline is the dilution fluid of choice. Other solutions may be approved by the Transfusion Facility (ABO compatible plasma, 5% albumin and plasma protein fraction). - Do NOT add Calcium or Dextrose-containing solutions or drugs to a component or infuse in the same IV line as the blood component because they may cause hemolysis or clotting. - Use of infusion pumps or blood warmers may require special tubing. - Transfusion tubing should be used as defined by the vendor in their manufacturer’s instructions. |
| --- | --- |

| ##### Reflex Testing | Results obtained when performing ordered tests may indicate the need for additional testing. The descriptions below are examples of reflex testing that may occur. - If an ABO typing discrepancy is detected, it may be resolved by special antigen typing, antibody identification, or other methods. - If the patient is Rh(D)-negative at immediate spin, weak D reflex testing may occur if the: - Patient history is D-positive - Patient is a D-negative infant from a D-negative female - Autologous unit tests D-positive - Detection of a positive antibody screen may lead to the performance of one or more antibody identification panels. Depending on the outcome of that testing, additional testing may be performed to characterize the patient’s antibody(ies) including the following: - Treatment of red cells with enzymes or other reagents - Treatment of serum or plasma by dilution or chemical modification - Adsorption of patient serum or plasma with autologous or allogeneic red cells, other reagent red cells, or red cell stroma. - If a new antibody specificity is identified, special antigen typing of the patient’s red blood cells for each corresponding antigen will be performed when possible. - If an autocontrol (autologous control cells tested along with antibody identification) is found to be reactive, one or more Direct Antiglobulin Tests (DAT) will be performed. - A positive DAT or autocontrol may result in preparation of an eluate from the patient’s red cells and subsequent antibody identification. - If antibody identification is ordered on a patient, and initial testing is negative, no further testing is normally performed. - Antibodies detected in a prenatal patient may result in titration of the patient’s serum or plasma for each clinically significant antibody. If an antibody titer is ordered and antibody identification has not been performed recently, antibody identification will be performed first. - When red blood cell products are ordered, patients may require different levels of crossmatch testing (electronic, serologic, or extended). - If the patient has clinically significant antibodies or has a history of such antibodies, red blood cell units will be tested for the corresponding antigen and an antiglobulin crossmatch will be performed. |
| --- | --- |

| ##### Critical Results | When a critical value has been obtained, the Transfusion Service will notify the Transfusion Facility as soon as possible. - If the Transfusion Facility is closed, the Transfusion Service staff will attempt to contact the Transfusion Facility physician, or the person designated as being on-call for the physician. - If the Transfusion Facility staff cannot be reached, the Transfusion Service (TS) staff will notify the TS Medical Director. Critical Results: - A transfusion reaction investigation that suggests any evidence of immunohematologic incompatibility, clerical error, blood administration error or specimen collection error - Incompatible crossmatch completed after emergency release of uncrossmatched red cells - Evidence of a delayed hemolytic transfusion reaction - Evidence of a transfusion error - ABO discrepancy of current type and screen specimen with previous record, not accounted for by known history - OB patient with new clinically significant antibody identified during pregnancy - Antibody titer in OB patient with clinically significant antibody or an increase in titer of more than 2 dilutions from a previously collected specimen - Positive DAT on cord blood, when the mother’s ABO/Rh is compatible with the infant’s ABO/Rh - Feto-maternal hemorrhage in excess of 30 mL whole blood - Delay of availability of blood components - Unavailable compatible blood components |
| --- | --- |

| ##### Type and Screen | - A Type and Screen may be ordered to obtain a current blood type and antibody screen result for a patient when there is a low likelihood for transfusion need but may not be needed for a surgical procedure. - Patient’s blood specimen is tested for: - ABO Group and Rh (D) Type - Antibody detection of unexpected antibodies - If there are no unexpected antibodies identified and the patient does not have a history or significant antibodies, the specimen is stored in the Transfusion Service for future blood product orders and crossmatching. - If transfusion becomes necessary, ABO/Rh compatible blood can be provided using the stored specimen. - If the patient has a positive antibody screen or a history of clinically significant antibodies, the Transfusion Service will contact the Transfusion Facility to determine the total number of components that may be needed. These units will be crossmatched for the patient in anticipation. |
| --- | --- |

| ##### Emergency Requirement for Blood | - Blood components may be provided for emergency use if required by the Transfusion Facility. This may include uncrossmatched group O red cell components as well as plasma, platelets, or cryoprecipitate. Specific requirements must be met by the Transfusion Facility including proper SOPs, oversight of the program, and use of specific forms. - The methods by which emergency blood components are provided vary with the capacity and need of the Transfusion Facility. Factors to consider include: - Acuity of patients treated at the Transfusion Facility - Capacity to provide properly monitored storage - Distance from the Transfusion Facility - Contact the Transfusion Service if emergency blood component service is required at the Transfusion Facility. |
| --- | --- |

Adverse Reactions

| ##### Signs and Symptoms of Transfusion Reactions | The following tables list the definitions and symptoms associated with types of reactions. A transfusion reaction represents any unfavorable event that may have a relationship to transfusion and should be investigated. Immediate Transfusion Reactions |
| --- | --- |

| Type of Reaction | Definition | Symptoms |
| --- | --- | --- |
| Transfusion associated sepsis | - Bacterial contamination of transfused blood | - Hypotension or Hypertension - Shaking chills - Hemoglobinuria - DIC - Oliguria/anuria - Fever - Tachycardia - Neutropenia (post transfusion) |
| Allergic Reaction | - Allergic reaction to transfusion which occurs during or within 2 hours of cessation of transfusion | - Conjunctival edema Edema of lips, tongue and uvula - Erythema and edema of periorbital area Generalized flushing Hypotention Localized angioedema Maculopapular rash Pruritis (itching) - Respiratory distress (bronchospasms) - Urticaria (hives) |

Continued on next page

Signs and Symptoms of Transfusion Reactions (continued)

| Type of Reaction | Definition | Symptoms |
| --- | --- | --- |
| Acute hemolytic transfusion reaction (AHTR) | - Transfused RBCs interact with pre-formed antibodies in recipient | - Back/flank pain - Chills/rigors - Discolored urine - Disseminated intravascular coagulation (DIC) - Epistaxis - Fever - Hypotension - Oligouria/anuria - Pain and/or oozing at IV site - Renal failure - Decreased fibrinogen - Decreased haptoglobin - Elevated bilirubin - Elevated LDH - Hemoglobinemia - Hemoglobinuria - Immune-mediated: - Positive DAT - Positive elution with alloantibody present - Non-immune mediated: - Thermal, osmotic, mechanical or chemical cause confirmed - Serologic testing is negative |
| Delayed hemolytic transfusion reaction (DHTR) | Development of new alloantibody in response to the transfused red cells. Symptoms/ signs of hemolysis more than 24 hours after transfusion. Can occur up to four weeks after the transfusion. | - Positive DAT 24 hours to 28 days after transfusion - Positive elution with alloantibody present - Newly identified alloantibody - Inadequate rise in hemoglobin post-transfusion or rapid - decrease in hemoglobin back to pre-transfusion levels - Unexplained appearance of spherocytes |
| Delayed serologic transfusion reaction (DSTR) | - New clinically significant alloantibody identified in laboratory serological testing. | - Positive antibody screening - Positive DAT |
| Febrile non-hemolytic reactions (FNHTR) | - Recipient’s HLA, granulocyte, and platelet-specific antibodies react with transfused antigens, and elicit the release of cytokines, that cause fever. | - Temperature increase > 1C or 2F, without any other explanation. - Chills - Rigors |

Continued on next page

Signs and Symptoms of Transfusion Reactions (continued)

| Type of Reaction | Definition | Symptoms |
| --- | --- | --- |
| Transfusion Associated Circulatory Overload (TACO) | - Acute pulmonary edema due to volume overload. | - Acute respiratory distress - Elevated brain natriuretic peptide (BPN) - Elevated central venous pressure (CVP) - Evidence of left heart failure - Evidence of positive fluid balance - Radiographic evidence of pulmonary edema |
| Transfusion-related acute lung injury (TRALI) | - A new episode of acute lung injury (ALI) that occurs during or within 6 hours of a completed transfusion. | - Acute respiratory insufficiency in the absence of evidence of circulatory overload. - No evidence of left atrial hypertension - Acute onset - Hypoxemia (oxygen saturation decreases to < 90% on room air) - Radiographic evidence of Bilateral infiltrates. - No other evidence of cardiac failure or reason for respiratory failure. - No evidence of acute lung injury prior to transfusion |
| Transfusion-associated dyspnea (TAD) | - Acute respiratory distress within 24 hours of transfusion. | - Ruled out: - Allergic reaction - TACO - TRALI |
| Hypotensive transfusion reaction | - Condition with low blood pressure associated with administration of blood products. The low blood pressure quickly resolves when the transfusion is stopped. All other adverse reactions with hypotension ruled out. | - Adults: - Drop in systolic BP > 30 mm Hg and < 80 mm Hg. - Infants and children: - > 25% drop in systolic BP from baseline. - Neonates: - > 25% drop in baseline value. |
| Transfusion-associated Graft-vs.-Host disease | - Immunologic complication caused by engraftment and proliferation of donor lymphocytes from a non-irradiated cellular blood component in a susceptible (immunocompromised) host. | - Symptom development 2 days to 6 weeks after transfusion. - Characteristic rash - Diarrhea - Fever - Hepatomegaly - Liver dysfunction - Pancytopenia |

Continued on next page

Signs and Symptoms of Transfusion Reactions (continued)

| Type of Reaction | Definition | Symptoms |
| --- | --- | --- |
| Post transfusion purpura (PTP) | - Development of an antibodies to the HPA-1 platelet antigen - Abrupt onset of severe thrombocytopenia an average of 9 days post transfusion (range 1-24 days) | - Precipitous fall in platelet count - Generalized purpura - Development of HPA-1a or another platelet specific alloantibody. |
| Transfusion transmitted infection (TTI) | - Evidence of pathogen (bacterial, viral, parasitic or other) in any of - the following. - Transfusion recipient - Donor of transfused product - Additional recipient of a product from the same donation. | - Evidence of the pathogen in: - The transfused component - The donor at the time of donation - An addition component from the same donation. - An additional recipient of a component from the same donation. |

| ##### Process for Immediate Adverse Reactions | In the event of a suspected transfusion reaction, discontinue the transfusion and evaluate the patient status. NOTE: Initiate transfusion reaction work up prior to releasing patient from medical care. If symptoms suggest a hemolytic or anaphylactic transfusion reaction, administer appropriate treatment and keep the patient under observation until Transfusion Service reports results of the work up. If patient symptoms are severe, consider transporting patient to an Emergency Center. |
| --- | --- |

| Step | Action |  |
| --- | --- | --- |
| 1 | Stop transfusion. - Do not disconnect the blood component. - Keep IV line open with slow infusion of saline. - Treatment of hypotension and promotion of adequate renal blood flow are primary concerns. - Avoid over-hydration. - Ensure adequate renal perfusion by monitoring measurement of urine output to achieve a rate above 100 mL/hour in adults, or as appropriate. |  |
| 2 | Notify the patient’s physician and the Transfusion Service immediately. |  |
| 3 | If | Then |
|  | A Transfusion Reaction Work up is ordered | Disconnect the blood component. - Examine the label on the blood container and patient identification to detect errors. - Document in patient’s chart. - Notify Transfusion Service. - Go to Step 4. |
|  | If, after evaluation of the patient, a transfusion reaction is not suspected, and a transfusion reaction is not ordered | - Continue with transfusion following physician’s orders. |
| 4 | Complete the Report of Transfusion Reaction, TS 009 or approved Transfusion Facility equivalent. |  |
| 5 | Draw one 7 mL EDTA and one 7 mL Red Top specimen carefully to avoid hemolysis. Label the specimens according to the Specimen Collection and Labeling section above. |  |
| 6 | If symptoms include hypotension and fever (for red cells or platelets), bacterial contamination is a possibility. - Draw a blood culture specimen immediately and send to your microbiology reference laboratory for culture. |  |

Continued on next page

Process for Immediate Adverse Reactions (continued)

| Step | Action |
| --- | --- |
| 7 | Submit blood component bag and attached infusion line/IV solutions, post reaction blood specimens to Transfusion Service, along with the Report of Transfusion Reaction, TS 009, or Electronic Medical Record system. |
| 8 | Transfusion Service notifies Facility and physician verbally, as soon as evidence of a hemolytic transfusion reaction is ruled out. |
| 9 | Upon completion of Workup, Transfusion Service notifies Facility and physician verbally, with follow-up written document of results. |

NOTE: If additional components are required subsequent to the Transfusion Reaction Work up, send a new patient specimen and a Transfusion Service Order to the Transfusion Service.

| ##### Process for Other Adverse Reactions | - Post transfusion, the recipient may develop symptoms related to other adverse reactions: - Delayed transfusion reaction - Transfusion transmitted diseases or associated infections - If symptoms develop, notify Transfusion Service. - Complete a Report of Transfusion Reaction, TS 009, or approved Transfusion Facility equivalent for symptoms relating to a delayed transfusion reaction or a transfusion-transmitted infection. - The forms and instructions for completing the forms are obtained from the Transfusion Service. |
| --- | --- |